## Cover
Document#: S0034
Revision: B
Project Start Date: 8/14/25
Project Name: Aurora PGA
Package: PGA Module
Test Hardware Description: Sonoma Air full size BIM
Hardware Design Originator: D.
Socket PN: Boyd TMA-12B-5-0060@A bridge and 702-0003244 socket
Positions per board: 5
Layer Count: 6-8
Included documnet File Names: TBD
Components: SMD
Board Thickness: .125 inch
Customer approval:
Name: ______________________________________
Title: _______________________________________
Signature: __________________________________
Date: _______________________________________

## Revision History
Revision History:
REV. | DATE | ORIGINATOR | DESCRIPTION OF CHANGE
A | 2025-08-14 | D | Initial Release

## HTOL Requirements (biasing, tem
Driving requirements: | 1. 9 channels current driver
2. 3 power rails 3V3 and N1V8 and negative
3. 1 signal driver 5Vpp + 4V bias
4. 1 thermistor read out
PIN_ID | Driving Requirement | Comment | Driving Configuration | In-Situ Monitoring | Notes |  |  |  | WEST | [22 empty] | EAST
A1 | N/C | [6 empty] | INDEX | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | Y
A2 | 700mA, DC, current drive |  | Current driver 1 signal |  | Connect to Current Source Daughter card #1 connector Pin B2 and A4 | expected less than 3V |  | 1 | A1 | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A3 | 700mA, DC return of A2 |  | Current driver 1 ground |  | GND |  |  | 2 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A4 | N/C | [6 empty] | 3 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A5 | N/C | [6 empty] | 4 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A6 | Return path of A7 |  | Signal generator 1 ground |  | GND |  |  | 5 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A7 | 5Vpp + 4V bias, 100MHz Square Wave |  | Signal generator 1 signal |  | Oscillator curcuit | See Oscillator Circuit Tab |  | 6 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A8 | Return path of A7 |  | Signal generator 1 ground |  | GND |  |  | 7 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C |  |  | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A9 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal | Yes | PS3 | HCPS VOUT 9 current sense INA214 see shunt resistor value to the right | Pin A9 is 3mA max out 6.8 ohm | 8 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A10 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal |  | PS3 | HCPS VOUT 9 |  | 9 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A11 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal |  | PS3 | HCPS VOUT 9 |  | 10 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A12 | N/C | [6 empty] | 11 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A13 | N/C | [6 empty] | 12 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A14 | N/C | [6 empty] | 13 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A15 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal | Yes | PS3 | HCPS VOUT 9 current sense INA214 see shunt resistor value to the right | Pin A15 is 400uA max out 50 ohm 1/8W | 14 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A16 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal |  | PS3 | HCPS VOUT 9 |  | 15 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A17 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal |  | PS3 | HCPS VOUT 9 |  | 16 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A18 | N/C | [6 empty] | 17 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A19 | N/C | [6 empty] | 18 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A20 | 1200mA, DC, current drive |  | Current driver 2 signal |  | Connect to Current Source Daughter card #2 connector Pin B2 and A4 | expected less than 3V |  | 19 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A21 | N/C | [6 empty] | 20 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A22 | 2400mA, DC, current drive |  | Current driver 3 signal |  | Connect to Current Source Daughter card #3 connector Pin B2 and A4 | expected less than 3V |  | 21 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A23 | 2400mA, DC, current drive |  | Current driver 4 signal |  | Connect to Current Source Daughter card #4 connector Pin B2 and A4 | expected less than 3V |  | 22 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A24 | 2400mA, DC, current drive |  | Current driver 5 signal |  | Connect to Current Source Daughter card #5 connector Pin B2 and A4 | expected less than 3V |  | 23 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A25 | N/C | [6 empty] | 24 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A26 | N/C | [6 empty] | 25 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A27 | N/C | [6 empty] | 26 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A28 | N/C | [6 empty] | 27 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A29 | N/C | [6 empty] | 28 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A30 | N/C | [6 empty] | 29 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A31 | N/C | [6 empty] | 30 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A32 | N/C | [6 empty] | 31 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A33 | N/C | [6 empty] | 32 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
A34 | N/C | [6 empty] | 33 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
B1 | N/C | [6 empty] | 34 | #NAME? | #NAME? | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | #NAME? | #NAME?
B2 | N/C | [6 empty] | 35 | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C
B3 | N/C | [6 empty] | 36 | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C | N/C
B4 | N/C | [6 empty] | 37 | N/C | N/C | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | N/C | N/C
B5 | N/C | [6 empty] | 38 | N/C | N/C | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | N/C | N/C
B6 | Return path of A7 |  | Signal generator 1 ground |  | GND |  |  | 39 | N/C | N/C | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | N/C | N/C
B7 | Return path of A7 |  | Signal generator 1 ground |  | GND |  |  | 40 | N/C | N/C | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | N/C | N/C
B8 | Return path of A7 |  | Signal generator 1 ground |  | GND
B9 | Return path of power supply 3 |  | Power supply 3 ground |  | GND |  |  |  | WEST C: SOUTH
B10 | Return path of power supply 3 |  | Power supply 3 ground |  | GND
B11 | Return path of power supply 3 |  | Power supply 3 ground |  | GND
B12 | N/C
B13 | N/C
B14 | N/C
B15 | Return path of power supply 3 |  | Power supply 3 ground |  | GND
B16 | Return path of power supply 3 |  | Power supply 3 ground |  | GND
B17 | Return path of power supply 3 |  | Power supply 3 ground |  | GND
B18 | N/C
B19 | Return path of A20, A22, A23, A24 |  | Current Driver 2,3,4,5 common ground |  | GND
B20 | Return path of A20, A22, A23, A24 |  | Current Driver 2,3,4,5 common ground |  | GND
B21 | Return path of A20, A22, A23, A24 |  | Current Driver 2,3,4,5 common ground |  | GND
B22 | Return path of A20, A22, A23, A24 |  | Current Driver 2,3,4,5 common ground |  | GND
B23 | Return path of A20, A22, A23, A24 |  | Current Driver 2,3,4,5 common ground |  | GND
B24 | Return path of A20, A22, A23, A24 |  | Current Driver 2,3,4,5 common ground |  | GND
B25 | N/C
B26 | N/C
B27 | N/C
B28 | N/C
B29 | N/C
B30 | N/C
B31 | N/C
B32 | N/C
B33 | N/C
B34 | N/C
C37 | Connect to 3.3V, DC, power rail, total expect 1A |  | Power supply 1 signal |  | PS1 | LCPS J6 1-5 | 1 Supply per DUT
C38 | Connect to 3.3V, DC, power rail, total expect 1A |  | Power supply 1 signal |  | PS1 | LCPS J6 1-5 | 1 Supply per DUT
C39 | Return path of C37, C38, C40, D37, U40, V37, V39, V40 |  | Power supply 1,2 common ground |  | GND
C40 | Connect to -1.2V, DC, negative power rail, expect < 50mA |  | Power supply 2 signal |  | PS2 | Negative supply | From HCPS VOUT 10
D37 | Connect to 3.3V, DC, power rail, total expect 1A |  | Power supply 1 signal |  | PS1 | LCPS J6 1-5 | 1 Supply per DUT
D38 | Return path of C37, C38, C40, D37, U40, V37, V39, V40 |  | Power supply 1,2 common ground |  | GND
D39 | Return path of C37, C38, C40, D37, U40, V37, V39, V40 |  | Power supply 1,2 common ground |  | GND
D40 | Return path of C37, C38, C40, D37, U40, V37, V39, V40 |  | Power supply 1,2 common ground |  | GND
E37 | N/C
E38 | N/C
E39 | N/C
E40 | N/C
F37 | N/C
F38 | N/C
F39 | N/C
F40 | N/C
G37 | N/C
G38 | N/C
G39 | N/C
G40 | N/C
H37 | N/C
H38 | N/C
H39 | N/C
H40 | N/C
I37 | N/C
I38 | N/C
I39 | N/C
I40 | N/C
J37 | N/C
J38 | N/C
J39 | N/C
J40 | N/C
K37 | N/C
K38 | N/C
K39 | N/C
K40 | N/C
L37 | N/C
L38 | N/C
L39 | N/C
L40 | N/C
M37 | N/C
M38 | N/C
M39 | N/C
M40 | N/C
N37 | N/C
N38 | N/C
N39 | N/C
N40 | N/C
O37 | N/C
O38 | N/C
O39 | N/C
O40 | N/C
P37 | N/C
P38 | N/C
P39 | N/C
P40 | N/C
Q37 | N/C
Q38 | N/C
Q39 | N/C
Q40 | N/C
R37 | N/C
R38 | N/C
R39 | N/C
R40 | N/C
S37 | N/C
S38 | N/C
S39 | N/C
S40 | N/C
T37 | N/C
T38 | N/C
T39 | N/C
T40 | N/C
U37 | Return path of C37, C38, C40, D37, U40, V37, V39, V40 |  | Power supply 1,2 common ground |  | GND
U38 | Return path of C37, C38, C40, D37, U40, V37, V39, V40 |  | Power supply 1,2 common ground |  | GND
U39 | Return path of C37, C38, C40, D37, U40, V37, V39, V40 |  | Power supply 1,2 common ground |  | GND
C30 | Connect to -18V, DC, negative power rail, expect < 50mA
U40 | 3.3V, DC, power rail, expect 1A |  | Power supply 1 signal |  | PS1 | LCPS J6 1-5 | 1 Supply per DUT
V37 | Connect to -1.2V, DC, negative power rail, expect < 50mA |  | Power supply 2 signal |  | PS2
V38 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal | Yes | PS3 | HCPS VOUT 9 current sense INA214 see shunt resistor value to the right | Pin V38 is 15mA max out 1.3 ohm 1W
V39 | 3.3V, DC, power rail, expect 1A |  | Power supply 1 signal |  | PS1 | LCPS J6 1-5 | 1 Supply per DUT
V40 | 3.3V, DC, power rail, expect 1A |  | Power supply 1 signal |  | PS1 | LCPS J6 1-5 | 1 Supply per DUT
W37 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal | Yes | PS3 | HCPS VOUT 9 current sense INA214 see shunt resistor value to the right | Pin W37 is 15mA max out 1.3 ohm 1W
W38 | Return path of power supply 3 |  | Power supply 3 ground |  | GND
W39 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal | Yes | PS3 | HCPS VOUT 9 current sense INA214 see shunt resistor value to the right | Pin W39 is 3mA max out 6.8 ohm 1/4W
W40 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal | Yes | PS3 | HCPS VOUT 9 current sense INA214 see shunt resistor value to the right | Pin W40 is 15mA max out 1.3 ohm 1W
X1 | Return path of power supply 3 |  | Power supply 3 ground |  | GND
X2 | Thermistor read out w/ Y2 |  | Thermistor read out | Yes | 1/16A fuse to 4.096VREF
X3 | N/C
X4 | N/C
X5 | N/C
X6 | N/C
X7 | N/C
X8 | N/C
X9 | N/C
X10 | N/C
X11 | N/C
X12 | N/C
X13 | Return path of Y13 |  | Current Driver 6 ground |  | GND
X14 | Thermistor read out w/ Y14 |  | Thermistor read out | Yes | 1/16A fuse to 4.096VREF
X15 | N/C
X16 | N/C
X17 | N/C
X18 | N/C
X19 | Thermistor read out w/ Y19 |  | Thermistor read out | Yes | 1/16A fuse to 4.096VREF
X20 | N/C
X21 | Return path of Y22, Y23, Y24 |  | Current Driver 7,8,9 common ground |  | GND
X22 | Return path of Y22, Y23, Y24 |  | Current Driver 7,8,9 common ground |  | GND
X23 | Return path of Y22, Y23, Y24 |  | Current Driver 7,8,9 common ground |  | GND
X24 | Return path of Y22, Y23, Y24 |  | Current Driver 7,8,9 common ground |  | GND
X25 | N/C
X26 | N/C
X27 | N/C
X28 | N/C
X29 | N/C
X30 | N/C
X31 | N/C
X32 | N/C
X33 | N/C
X34 | N/C
Y1 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal | Yes | PS3 | HCPS VOUT 9 current sense INA214 see shunt resistor value to the right | Pin Y1 is 10mA max out 2 ohm 1/2W
Y2 | Thermistor read out w/ X2 |  | Thermistor read out | Yes | connect to 3.74k 0.1% resistor to XADC and 1.24k 0.1% to GND See XADC ADC resource page
Y3 | N/C
Y4 | N/C
Y5 | N/C
Y6 | N/C
Y7 | N/C
Y8 | N/C
Y9 | N/C
Y10 | N/C
Y11 | N/C
Y12 | N/C
Y13 | 1200mA, DC, current drive |  | Current Driver 6 |  | Connect to Current Source Daughter card #6 connector Pin B2 and A4 | expected less than 3V
Y14 | Thermistor read out w/ X14 |  | Thermistor read out | Yes | connect to 3.74k 0.1% resistor to XADC and 1.24k 0.1% to GND See XADC ADC resource page
Y15 | N/C
Y16 | N/C
Y17 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal | Yes | PS3 | HCPS VOUT 9 current sense INA214 see shunt resistor value to the right | Pin Y17 is 6mA max out 3.3 ohm 1/2W
Y18 | Connect to 1.8V, DC, power rail, total expect <15mA |  | Power supply 3 signal |  | PS3 | HCPS VOUT 9
Y19 | Thermistor read out w/ X19 |  | Thermistor read out | Yes | connect to 3.74k 0.1% resistor to XADC and 1.24k 0.1% to GND See XADC ADC resource page
Y20 | N/C
Y21 | N/C
Y22 | 2400mA, DC, current drive |  | Current Driver 7 |  | Connect to Current Source Daughter card #7 connector Pin B2 and A4 | expected less than 3V
Y23 | 2400mA, DC, current drive |  | Current Driver 8 |  | Connect to Current Source Daughter card #8 connector Pin B2 and A4 | expected less than 3V
Y24 | 2400mA, DC, current drive |  | Current Driver 9 |  | Connect to Current Source Daughter card #9 connector Pin B2 and A4 | expected less than 3V
Y25 | N/C
Y26 | N/C
Y27 | N/C
Y28 | N/C
Y29 | N/C
Y30 | N/C
Y31 | N/C
Y32 | N/C
Y33 | N/C
Y34 | N/C

## Oscillator Circuit
In Connect to Sonoma High speed Channel | V2 PS1 3.3V J6 VOUT1-5 thru a fuse one supply per DUT.
V1 5V J6 VOUT8 one per DUT thru a fuse
Out to DUT | 12V connect to System 12V tru a fuse
OSCILLATOR IS DNI but order it for Assembly

## Daughter card Design (hidden)
use Sullins S9641-ND same as in S0024
A1 | PS4 | B1 | GND
A2 | AF36 | B2 | Y2
A3 | NC | B3 | NC
A4 | Y2 | B4 | AF36
A5 | GND | B5 | PS4
The card will use 2 LT3083 Ics TSSOP 16
IN connect to PS4 and 10uf to GND for each IC.
VCONTROL CONNECT TO Sullins B3 and to IN THRU 0 ohm
OUT IC1 connect to Sullins B2, A4 thru 0 ohm and 10uf to GND
OUT IC2 connect to Sullins A2, B4 thru 0 ohm and 10uf to GND
SET connect to OUT thru 665 ohm one oper IC

## Daughter card pinout
use Sullins S9641-ND same footprint as in B1830 or S0033
|  |  | Card 1-5 | Card 6-9
| IN comes from LCPS | DUT1 | LCPS6-VOUT6 | LCPS6-VOUT7
|  | DUT2 | LCPS7-1 | LCPS7-2
|  | DUT3 | LCPS7-3 | LCPS7-4
|  | DUT4 | LCPS7-5 | LCPS7-6
|  | DUT5 | LCPS7-7 | LCPS7-8

## Current source (hidden)
LM317L

## Pinlist (hidden)
X107 die pad # | order | X107 die pad name | P or N | Bond to aQFN lead | Pad to Pad bond | aQFN Lead Number | Burn-in connections | New HTOL connections | Voltage (V) | Current (mA) | Board Power Key | What to Monitor | Direction | Voltage (V) | Current (mA) | Board Power Key | What to monitor
9 | 9 | THERM | P | Yes | No | Y2 | 20K to 1V and 0 ohm to Monitor channel | DNI resistor to GND 0 ohm resistor to Monitor channel MUX and 1/2K resistor to PS1 1.4V | 1V | 1mA | PS1=1V | Monitor Voltage (provision for all DUT, but insitu measurement on 14 DUTs out of 25) | PS1=1V, pin 9 to 8 | 1V | 1mA | PS1=1V | Monitor Voltage (provision for all DUT, but insitu measurement on 14 DUTs out of 25)
60 | 60 | VTERM1 |  | Yes | No | E39 | 50mA, 36ohm, connect to PS3=1.8V | 32.5ohm, connect to GND |  | 50mA | PS3=1.8V |  | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V
61 | 61 | VTERM2 |  | Yes | No | C41 | 50mA, 36ohm, connect to PS3=1.8V | 32.5ohm, connect to GND |  | 50mA | PS3=1.8V |  | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V
62 | 62 | VTERM3 |  | Yes | No | C39 | 50mA, 36ohm, connect to PS3=1.8V | 32.5ohm, connect to GND |  | 50mA | PS3=1.8V |  | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V
63 | 63 | VTERM4 |  | Yes | No | A39 | 50mA, 36ohm, connect to PS3=1.8V | 32.5ohm, connect to GND |  | 50mA | PS3=1.8V |  | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V
11 | 11 | VOA_LOY | P | Yes | No | AB2 | Apply PS2=1.4V on Pin11 | 31 ohm to PS1 1.4V | 1.4V | 45mA | PS2=1.4V |  | Apply PS2=1.4V on Pin11 | 1.4V | 45mA | PS2=1.4V
12 | 12 | VOA_LOX | P | Yes | No | AD2 | Apply PS2=1.4V on Pin12 | 31 ohm to PS1 1.4V | 1.4V | 45mA | PS2=1.4V |  | Apply PS2=1.4V on Pin12 | 1.4V | 45mA | PS2=1.4V
14 | 14 | VPD_YQ | N | Yes | No | AE5 | Apply PS5= +2V to Pin13 | to regulator from PS3 4V output of regulator is V5 2V one per board | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage | Apply PS5= +2V to Pin13 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage
17 | 17 | VPD_YI | N | Yes | No | AF8 | Apply PS5= +2V to Pin17 | to regulator from PS3 4V output of regulator is V5 2V one per board | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage | Apply PS5= +2V to Pin17 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage
20 | 20 | VPD_XQ | N | Yes | No | AG11 | Apply PS5= +2V to Pin20 | to regulator from PS3 4V output of regulator is V5 2V one per board | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage | Apply PS5= +2V to Pin20 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage
50 | 50 | VPM | N | Yes | No | M40 | Apply PS5= +2V to Pin50 | to regulator from PS3 4V output of regulator is V5 2V one per board | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage | Apply PS5= +2V to Pin50 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage
23 | 23 | VPD_XI | N | Yes | No | AG13 | Apply PS6= +4V to Pin23 | to PS3 4V thru 32.5 ohm and 0 ohm to Monitor channel MUX one per DUT | +4V | 1mA | PS6= +4V | Monitor Voltage HS PD at 4V | Apply PS6= +4V to Pin23 | +4V | 1mA | PS6= +4V | Monitor Voltage HS PD at 4V
7 | 7 | RXPD_VPM | N | Yes | No | V2 | Apply PS6= +4V to Pin7 | to PS3 4V thru 32.5 ohm and 0 ohm to Monitor channel MUX one per DUT | PS6= +4V | 1mA | PS6= +4V | Monitor Voltage MPD at 4V | Apply PS6= +4V to Pin7 | PS6= +4V | 1mA | PS6= +4V | Monitor Voltage MPD at 4V
59 | 59 | VOA_XI | P | Yes | No | F40 | Appy PS2=1.4V on Pin59 | 31 ohm to PS1 1.4V | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin59 | 1.4V | 45mA | PS2=1.4V
65 | 65 | VOA_XQ | P | Yes | No | A37 | Appy PS2=1.4V on Pin65 | 31 ohm to PS1 1.4V | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin65 | 1.4V | 45mA | PS2=1.4V
70 | 70 | VOA_X | P | Yes | No | B32 | Appy PS2=1.4V on Pin70 | 31 ohm to PS1 1.4V | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin70 | 1.4V | 45mA | PS2=1.4V
76 | 76 | VOA_YI | P | Yes | No | A27 | Appy PS2=1.4V on Pin76 | 31 ohm to PS1 1.4V | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin76 | 1.4V | 45mA | PS2=1.4V
79 | 79 | VOA_YQ | P | Yes | No | A25 | Appy PS2=1.4V on Pin79 | 31 ohm to PS1 1.4V | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin79 | 1.4V | 45mA | PS2=1.4V
83 | 83 | VOA_Y | P | Yes | No | C21 | Appy PS2=1.4V on Pin83 | 31 ohm to PS1 1.4V | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin83 | 1.4V | 45mA | PS2=1.4V
45 | 45 | BIAS_IN2 |  | Yes | No | R39 | Floating | NC |  |  |  |  | Floating
46 | 46 | BIAS_IN1 |  | Yes | No | R41 | Floating | NC |  |  |  |  | Floating
47 | 47 | BIAS_QN2 |  | Yes | No | P40 | Floating | NC |  |  |  |  | Floating
48 | 48 | BIAS_QN1 |  | Yes | No | N39 | Floating | NC |  |  |  |  | Floating
49 | 49 | MZBIAS |  | Yes | No | N41 | Floating | NC |  |  |  |  | Floating
1 | 1 | LO_XPM | P | Yes | No | P2 | GND | GND |  |  |  |  | GND
2 | 2 | LO_YPM | P | Yes | No | R1 | GND | GND |  |  |  |  | GND
5 | 5 | RXPD_X | P | Yes | No | T2 | GND | GND |  |  |  |  | GND
6 | 6 | RXPD_Y | P | Yes | No | U1 | GND | GND |  |  |  |  | GND
8 | 8 | TNERMGND | N | Yes | No | W1 | GND | GND |  |  |  |  | Temp Sense GND
10 | 10 | VOA_BIAS | N | Yes | No | AA1 | GND | GND |  |  |  |  | GND
13 | 13 | RX_YQA | P | Yes | No | AE3 | GND | GND |  |  |  |  | GND
15 | 15 | RX_YQB | P | Yes | No | AG5 | GND | GND |  |  |  |  | GND
16 | 16 | RX_YIA | P | Yes | No | AE7 | GND | GND |  |  |  |  | GND
18 | 18 | RX_YIB | P | Yes | No | AE9 | GND | GND |  |  |  |  | GND
19 | 19 | RX_XQA | P | Yes | No | AF10 | GND | GND |  |  |  |  | GND
21 | 21 | RX_XQB | P | Yes | No | AE11 | GND | GND |  |  |  |  | GND
22 | 22 | RX_XIA | P | Yes | No | AF12 | GND | GND |  |  |  |  | GND
24 | 24 | RX_XIB | P | Yes | No | AE15 | GND | GND |  |  |  |  | GND
25 | 25 | GNDYQ1 | P | Yes | No | AE23 | GND | GND |  |  |  |  | GND
27 | 27 | GNDYQ2 | P | Yes | No | AE25 | GND | GND |  |  |  |  | GND
29 | 29 | GNDYQ3 | P | Yes | No | AF26 | GND | GND |  |  |  |  | GND
30 | 30 | GNDYI1 | P |  |  | AE27 | GND | GND |  |  |  |  | GND |  |  |  |  | 1 tHerm needs 1V
32 | 32 | GNDYI2 | P | Yes | No | AF28 | GND | GND |  |  |  |  | GND |  |  |  |  | 2 All 1.4V pins are current source pins
34 | 34 | GNDYI3 | P | Yes | No | AG29 | GND | GND |  |  |  |  | GND |  |  |  |  | 3 remove active current monitor on 1.4V
35 | 35 | GNDXQ1 | P |  |  | AF30 | GND | GND |  |  |  |  | GND
37 | 37 | GNDXQ2 | P | Yes | No | AE31 | GND | GND |  |  |  |  | GND
39 | 39 | GNDXQ3 | P | Yes | No | AG33 | GND | GND |  |  |  |  | GND
40 | 40 | GNDXI1 | P |  |  | AE33 | GND | GND |  |  |  |  | GND
42 | 42 | GNDXI2 | P | Yes | No | AG35 | GND | GND |  |  |  |  | GND
44 | 44 | GNDXI3 | P | Yes | No | AF36 | GND | GND |  |  |  |  | GND
51 | 51 | PM_XI | P | Yes | No | L39 | GND | GND |  |  |  |  | GND
52 | 52 | PHASEGND_XI |  | Yes | No | L41 | GND | GND |  |  |  |  | GND
55 | 55 | PM_XQ | P | Yes | No | J39 | GND | GND |  |  |  |  | GND
56 | 56 | PM_X | P | Yes | No | H40 | GND | GND |  |  |  |  | GND
57 | 57 | PHASEGND_X |  | Yes | No | G41 | GND | GND |  |  |  |  | GND
64 | 64 | VOAGND_XIQ | N | Yes | No | C37 | GND | GND |  |  |  |  | GND
66 | 66 | PHASEGND_XQ |  | Yes | No | C35 | GND | GND |  |  |  |  | GND
69 | 69 | PM_YI | P | Yes | No | C33 | GND | GND |  |  |  |  | GND
73 | 73 | PHASEGND_Y |  | Yes | No | C29 | GND | GND |  |  |  |  | GND
75 | 75 | YI_TAP | P | Yes | No | B28 | GND | GND |  |  |  |  | GND
77 | 77 | PM_Y | P | Yes | No | C27 | GND | GND |  |  |  |  | GND
78 | 78 | PM_YQ | P | Yes | No | B26 | GND | GND |  |  |  |  | GND
80 | 80 | VOAGND_YIQ | N | Yes | No | C25 | GND | GND |  |  |  |  | GND
81 | 81 | PHASEGND_YQ |  | Yes | No | A23 | GND | GND |  |  |  |  | GND
84 | 84 | VOAGND_XY | N | Yes | No | A33 | GND | GND |  |  |  |  | GND
85 | 85 | PHASEGND_YI |  | Yes | No | A31 | GND | GND |  |  |  |  | GND
86 | 86 | YQ_TAP | P | Yes | No | A21 | GND | GND |  |  |  |  | GND
87 | 87 | TAP_Y | P | Yes | No | C19 | GND | GND |  |  |  |  | GND
54 | 54 | XI_TAP | P | Yes | No | J41 | GND | GND |  |  |  |  | GND
68 | 68 | XQ_TAP | P | Yes | No | B34 | GND | GND |  |  |  |  | GND
71 | 71 | TAP_X | P | Yes | No | C31 | GND | GND |  |  |  |  | GND
3 | 3 | PHASE_SPLITTER |  | No | No | NC | NC | NC |  |  |  |  | Floating
4 | 4 | PHASEGND_SPLITTER |  | No | No | NC | NC | NC |  |  |  |  | Floating
53 | 53 | PHASE_XI |  | Yes | No | K40 | PS3=1.8V, Pin 53 to 52(GND) | to regulator from PS3 4V output of regulator is V6 1.8V oneregulator per 20 parts | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 53 to 52(GND) | 1.8V | 24mA | PS3=1.8V
58 | 58 | Phase_X |  | Yes | No | G39 | PS3=1.8V, Pin 58 to 57(GND) | to regulator from PS3 4V output of regulator is V6 1.8V oneregulator per 20 parts | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 58 to 57(GND) | 1.8V | 24mA | PS3=1.8V
67 | 67 | PHASE_XQ |  | Yes | No | A35 | PS3=1.8V, Pin 67 to 66(GND) | to regulator from PS3 4V output of regulator is V6 1.8V oneregulator per 20 parts | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 67 to 66(GND) | 1.8V | 24mA | PS3=1.8V
72 | 72 | PHASE_YI |  | Yes | No | B30 | PS3=1.8V, Pin 72 to 85(GND) | to regulator from PS3 4V output of regulator is V6 1.8V oneregulator per 20 parts | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 72 to 85(GND) | 1.8V | 24mA | PS3=1.8V
74 | 74 | PHASE_Y |  | Yes | No | A29 | PS3=1.8V, Pin 74 to 73(GND) | to regulator from PS3 4V output of regulator is V6 1.8V oneregulator per 20 parts | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 74 to 73(GND) | 1.8V | 24mA | PS3=1.8V
82 | 82 | PHASE_YQ |  | Yes | No | C23 | PS3=1.8V, Pin 82 to 81(GND) | to regulator from PS3 4V output of regulator is V6 1.8V oneregulator per 20 parts | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 82 to 81(GND) | 1.8V | 24mA | PS3=1.8V
26 | 26 | RFYQA | N | Yes | No | AF24 | PS7= +5V, pin 26 to 27(GND) | to PS4 5V | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 26 to 27(GND) | +5V | 0.003mA | PS7= +5V
28 | 28 | RFYQB | N | Yes | No | AG25 | PS7= +5V, pin 28 to 29(GND) | to PS4 5V | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 28 to 29(GND) | +5V | 0.003mA | PS7= +5V
31 | 31 | RFYIA | N | Yes | No | AG27 | PS7= +5V, pin 31 to 32(GND) | to PS4 5V | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 31 to 32(GND) | +5V | 0.003mA | PS7= +5V
33 | 33 | RFYIB | N | Yes | No | AE29 | PS7= +5V, pin 33 to 34(GND) | to PS4 5V | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 33 to 34(GND) | +5V | 0.003mA | PS7= +5V
36 | 36 | RFXQA | N | Yes | No | AG31 | PS7= +5V, pin 36 to 37(GND) | to PS4 5V | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 36 to 37(GND) | +5V | 0.003mA | PS7= +5V
38 | 38 | RFXQB | N | Yes | No | AF32 | PS7= +5V, pin 38 to 39(GND) | to PS4 5V | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 38 to 39(GND) | +5V | 0.003mA | PS7= +5V
41 | 41 | RFXIA | N | Yes | No | AF34 | PS7= +5V, pin 41 to 42(GND) | to PS4 5V | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 41 to 42(GND) | +5V | 0.003mA | PS7= +5V
43 | 43 | RFXIB | N | Yes | No | AE35 | PS7= +5V, pin 43 to 44(GND) | to PS4 5V | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 43 to 44(GND) | +5V | 0.003mA | PS7= +5V
88 | 88 | VTERM5 |  | No | Yes | D1.94 |  | NC
89 | 89 | VTERM6 |  | No | Yes | D1.95 |  | NC
90 | 90 | VTERM7 |  | No | Yes | D1.96 |  | NC
91 | 91 | VTERM8 |  | No | Yes | D1.92 |  | NC
92 | 92 | VTERM9 |  | No | Yes | D1.91 |  | NC
93 | 93 | VTERM10 |  | No | Yes | D1.97 |  | NC
94 | 94 | VTERM_XI |  | No | Yes | D1.88 |  | NC
95 | 95 | VTERM_XQ |  | No | Yes | D1.89 |  | NC
96 | 96 | VTERM_YI |  | No | Yes | D1.90 |  | NC
97 | 97 | VTERM_YQ |  | No | Yes | D1.93 |  | NC

## X107-AAH@A267480006-A (hidden)
CONFIDENTIAL - DISCLOSED UNDER NDA
[6 empty] | Board#1 Bias Conditions (125°C) |  |  |  |  | Board#2 Bias Conditions (150°C)
X107 die pad # | X107 die pad name | P or N | Bond to aQFN lead | Pad to Pad bond | aQFN Lead Number | Direction | Voltage (V) | Current (mA) | Board Power Key | What to Monitor | Direction | Voltage (V) | Current (mA) | Board Power Key | What to monitor
1 | LO_XPM | P | Yes | No | P2 | GND |  |  |  |  | GND
2 | LO_YPM | P | Yes | No | R1 | GND |  |  |  |  | GND
3 | PHASE_SPLITTER |  | No | No | NC | Floating |  |  |  |  | Floating
4 | PHASEGND_SPLITTER |  | No | No | NC | Floating |  |  |  |  | Floating
5 | RXPD_X | P | Yes | No | T2 | GND |  |  |  |  | GND
6 | RXPD_Y | P | Yes | No | U1 | GND |  |  |  |  | GND
7 | RXPD_VPM | N | Yes | No | V2 | Apply PS6= +4V to Pin7 | PS6= +4V | 1mA | PS6= +4V | Monitor Voltage MPD at 4V | Apply PS6= +4V to Pin7 | PS6= +4V | 1mA | PS6= +4V | Monitor Voltage MPD at 4V
8 | TNERMGND | N | Yes | No | W1 | Temp Sense GND |  |  |  |  | Temp Sense GND
9 | THERM | P | Yes | No | Y2 | PS1=1V, pin 9 to 8 | 1V | 1mA | PS1=1V | Monitor Voltage (provision for all DUT, but insitu measurement on 14 DUTs out of 25) | PS1=1V, pin 9 to 8 | 1V | 1mA | PS1=1V | Monitor Voltage (provision for all DUT, but insitu measurement on 14 DUTs out of 25)
10 | VOA_BIAS | N | Yes | No | AA1 | GND |  |  |  |  | GND
11 | VOA_LOY | P | Yes | No | AB2 | Apply PS2=1.4V on Pin11 | 1.4V | 45mA | PS2=1.4V |  | Apply PS2=1.4V on Pin11 | 1.4V | 45mA | PS2=1.4V
12 | VOA_LOX | P | Yes | No | AD2 | Apply PS2=1.4V on Pin12 | 1.4V | 45mA | PS2=1.4V |  | Apply PS2=1.4V on Pin12 | 1.4V | 45mA | PS2=1.4V
13 | RX_YQA | P | Yes | No | AE3 | GND |  |  |  |  | GND
14 | VPD_YQ | N | Yes | No | AE5 | Apply PS5= +2V to Pin13 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage | Apply PS5= +2V to Pin13 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage
15 | RX_YQB | P | Yes | No | AG5 | GND |  |  |  |  | GND
16 | RX_YIA | P | Yes | No | AE7 | GND |  |  |  |  | GND
17 | VPD_YI | N | Yes | No | AF8 | Apply PS5= +2V to Pin17 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage | Apply PS5= +2V to Pin17 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage
18 | RX_YIB | P | Yes | No | AE9 | GND |  |  |  |  | GND
19 | RX_XQA | P | Yes | No | AF10 | GND |  |  |  |  | GND
20 | VPD_XQ | N | Yes | No | AG11 | Apply PS5= +2V to Pin20 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage | Apply PS5= +2V to Pin20 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage
21 | RX_XQB | P | Yes | No | AE11 | GND |  |  |  |  | GND
22 | RX_XIA | P | Yes | No | AF12 | GND |  |  |  |  | GND
23 | VPD_XI | N | Yes | No | AG13 | Apply PS6= +4V to Pin23 | +4V | 1mA | PS6= +4V | Monitor Voltage HS PD at 4V | Apply PS6= +4V to Pin23 | +4V | 1mA | PS6= +4V | Monitor Voltage HS PD at 4V
24 | RX_XIB | P | Yes | No | AE15 | GND |  |  |  |  | GND
25 | GNDYQ1 | P | Yes | No | AE23 | GND |  |  |  |  | GND
26 | RFYQA | N | Yes | No | AF24 | PS7= +5V, pin 26 to 27(GND) | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 26 to 27(GND) | +5V | 0.003mA | PS7= +5V
27 | GNDYQ2 | P | Yes | No | AE25 | GND |  |  |  |  | GND
28 | RFYQB | N | Yes | No | AG25 | PS7= +5V, pin 28 to 29(GND) | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 28 to 29(GND) | +5V | 0.003mA | PS7= +5V
29 | GNDYQ3 | P | Yes | No | AF26 | GND |  |  |  |  | GND
30 | GNDYI1 | P |  |  | AE27 | GND |  |  |  |  | GND
31 | RFYIA | N | Yes | No | AG27 | PS7= +5V, pin 31 to 32(GND) | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 31 to 32(GND) | +5V | 0.003mA | PS7= +5V
32 | GNDYI2 | P | Yes | No | AF28 | GND |  |  |  |  | GND
33 | RFYIB | N | Yes | No | AE29 | PS7= +5V, pin 33 to 34(GND) | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 33 to 34(GND) | +5V | 0.003mA | PS7= +5V
34 | GNDYI3 | P | Yes | No | AG29 | GND |  |  |  |  | GND
35 | GNDXQ1 | P |  |  | AF30 | GND |  |  |  |  | GND
36 | RFXQA | N | Yes | No | AG31 | PS7= +5V, pin 36 to 37(GND) | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 36 to 37(GND) | +5V | 0.003mA | PS7= +5V
37 | GNDXQ2 | P | Yes | No | AE31 | GND |  |  |  |  | GND
38 | RFXQB | N | Yes | No | AF32 | PS7= +5V, pin 38 to 39(GND) | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 38 to 39(GND) | +5V | 0.003mA | PS7= +5V
39 | GNDXQ3 | P | Yes | No | AG33 | GND |  |  |  |  | GND
40 | GNDXI1 | P |  |  | AE33 | GND |  |  |  |  | GND
41 | RFXIA | N | Yes | No | AF34 | PS7= +5V, pin 41 to 42(GND) | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 41 to 42(GND) | +5V | 0.003mA | PS7= +5V
42 | GNDXI2 | P | Yes | No | AG35 | GND |  |  |  |  | GND
43 | RFXIB | N | Yes | No | AE35 | PS7= +5V, pin 43 to 44(GND) | +5V | 0.003mA | PS7= +5V |  | PS7= +5V, pin 43 to 44(GND) | +5V | 0.003mA | PS7= +5V
44 | GNDXI3 | P | Yes | No | AF36 | GND |  |  |  |  | GND
45 | BIAS_IN2 |  | Yes | No | R39 | Floating |  |  |  |  | Floating
46 | BIAS_IN1 |  | Yes | No | R41 | Floating |  |  |  |  | Floating
47 | BIAS_QN2 |  | Yes | No | P40 | Floating |  |  |  |  | Floating
48 | BIAS_QN1 |  | Yes | No | N39 | Floating |  |  |  |  | Floating
49 | MZBIAS |  | Yes | No | N41 | Floating |  |  |  |  | Floating
50 | VPM | N | Yes | No | M40 | Apply PS5= +2V to Pin50 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage | Apply PS5= +2V to Pin50 | +2V | 1mA | PS5=+2V | Provision for Monitor Voltage
51 | PM_XI | P | Yes | No | L39 | GND |  |  |  |  | GND
52 | PHASEGND_XI |  | Yes | No | L41 | GND |  |  |  |  | GND
53 | PHASE_XI |  | Yes | No | K40 | PS3=1.8V, Pin 53 to 52(GND) | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 53 to 52(GND) | 1.8V | 24mA | PS3=1.8V
54 | XI_TAP | P | Yes | No | J41 | GND |  |  |  |  | GND
55 | PM_XQ | P | Yes | No | J39 | GND |  |  |  |  | GND
56 | PM_X | P | Yes | No | H40 | GND |  |  |  |  | GND
57 | PHASEGND_X |  | Yes | No | G41 | GND |  |  |  |  | GND
58 | Phase_X |  | Yes | No | G39 | PS3=1.8V, Pin 58 to 57(GND) | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 58 to 57(GND) | 1.8V | 24mA | PS3=1.8V
59 | VOA_XI | P | Yes | No | F40 | Appy PS2=1.4V on Pin59 | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin59 | 1.4V | 45mA | PS2=1.4V
60 | VTERM1 |  | Yes | No | E39 | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V |  | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V
61 | VTERM2 |  | Yes | No | C41 | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V |  | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V
62 | VTERM3 |  | Yes | No | C39 | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V |  | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V
63 | VTERM4 |  | Yes | No | A39 | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V |  | 50mA, 32.5ohm, connect to PS3=1.6V |  | 50mA | PS3=1.6V
64 | VOAGND_XIQ | N | Yes | No | C37 | GND |  |  |  |  | GND
65 | VOA_XQ | P | Yes | No | A37 | Appy PS2=1.4V on Pin65 | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin65 | 1.4V | 45mA | PS2=1.4V
66 | PHASEGND_XQ |  | Yes | No | C35 | GND |  |  |  |  | GND
67 | PHASE_XQ |  | Yes | No | A35 | PS3=1.8V, Pin 67 to 66(GND) | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 67 to 66(GND) | 1.8V | 24mA | PS3=1.8V
68 | XQ_TAP | P | Yes | No | B34 | GND |  |  |  |  | GND
69 | PM_YI | P | Yes | No | C33 | GND |  |  |  |  | GND
70 | VOA_X | P | Yes | No | B32 | Appy PS2=1.4V on Pin70 | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin70 | 1.4V | 45mA | PS2=1.4V
71 | TAP_X | P | Yes | No | C31 | GND |  |  |  |  | GND
72 | PHASE_YI |  | Yes | No | B30 | PS3=1.8V, Pin 72 to 85(GND) | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 72 to 85(GND) | 1.8V | 24mA | PS3=1.8V
73 | PHASEGND_Y |  | Yes | No | C29 | GND |  |  |  |  | GND
74 | PHASE_Y |  | Yes | No | A29 | PS3=1.8V, Pin 74 to 73(GND) | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 74 to 73(GND) | 1.8V | 24mA | PS3=1.8V
75 | YI_TAP | P | Yes | No | B28 | GND |  |  |  |  | GND
76 | VOA_YI | P | Yes | No | A27 | Appy PS2=1.4V on Pin76 | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin76 | 1.4V | 45mA | PS2=1.4V
77 | PM_Y | P | Yes | No | C27 | GND |  |  |  |  | GND
78 | PM_YQ | P | Yes | No | B26 | GND |  |  |  |  | GND
79 | VOA_YQ | P | Yes | No | A25 | Appy PS2=1.4V on Pin79 | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin79 | 1.4V | 45mA | PS2=1.4V
80 | VOAGND_YIQ | N | Yes | No | C25 | GND |  |  |  |  | GND
81 | PHASEGND_YQ |  | Yes | No | A23 | GND |  |  |  |  | GND
82 | PHASE_YQ |  | Yes | No | C23 | PS3=1.8V, Pin 82 to 81(GND) | 1.8V | 24mA | PS3=1.8V |  | PS3=1.8V, Pin 82 to 81(GND) | 1.8V | 24mA | PS3=1.8V
83 | VOA_Y | P | Yes | No | C21 | Appy PS2=1.4V on Pin83 | 1.4V | 45mA | PS2=1.4V |  | Appy PS2=1.4V on Pin83 | 1.4V | 45mA | PS2=1.4V
84 | VOAGND_XY | N | Yes | No | A33 | GND |  |  |  |  | GND
85 | PHASEGND_YI |  | Yes | No | A31 | GND |  |  |  |  | GND
86 | YQ_TAP | P | Yes | No | A21 | GND |  |  |  |  | GND
87 | TAP_Y | P | Yes | No | C19 | GND |  |  |  |  | GND
88 | VTERM5 |  | No | Yes | D1.94
89 | VTERM6 |  | No | Yes | D1.95
90 | VTERM7 |  | No | Yes | D1.96
91 | VTERM8 |  | No | Yes | D1.92
92 | VTERM9 |  | No | Yes | D1.91
93 | VTERM10 |  | No | Yes | D1.97
94 | VTERM_XI |  | No | Yes | D1.88
95 | VTERM_XQ |  | No | Yes | D1.89
96 | VTERM_YI |  | No | Yes | D1.90
97 | VTERM_YQ |  | No | Yes | D1.93

## IN5634SE Pinlist (hidden)
Pin Name | Pin No. | Pin Description | Function | Pin Type | Pin setup for HTOL
DINN_1 | 1 | Channel 1 data input (inverting) | Input | RF Input | 50 Mhz input from oscillator bufer N (350-500mV Vdiff) See supporting circuit page
DINN_2 | 3 | Channel 2 data input (inverting) | Input | RF Input | 50 Mhz input from oscillator bufer N (350-500mV Vdiff) See supporting circuit page
DINN_3 | 5 | Channel 3 data input (inverting) | Input | RF Input | 50 Mhz input from oscillator bufer N (350-500mV Vdiff) See supporting circuit page
DINN_4 | 7 | Channel 4 data input (inverting) | Input | RF Input | 50 Mhz input from oscillator bufer N (350-500mV Vdiff) See supporting circuit page
DINP_1 | 2 | Channel 1 data input (non-inverting) | Input | RF Input | 50 Mhz input from oscillator bufer P (350-500mV Vdiff) See supporting circuit page
DINP_2 | 4 | Channel 2 data input (non-inverting) | Input | RF Input | 50 Mhz input from oscillator bufer P (350-500mV Vdiff) See supporting circuit page
DINP_3 | 6 | Channel 3 data input (non-inverting) | Input | RF Input | 50 Mhz input from oscillator bufer P (350-500mV Vdiff) See supporting circuit page
DINP_4 | 8 | Channel 4 data input (non-inverting) | Input | RF Input | 50 Mhz input from oscillator bufer P (350-500mV Vdiff) See supporting circuit page
GND | Center pad | Ground | Ground | Ground | GND
OUT1 | 23 | Channel 1 output | Output | RF Output | 50 ohm 1W to GND
OUT2 | 22 | Channel 2 output | Output | RF Output | 50 ohm 1W to GND
OUT3 | 21 | Channel 3 output | Output | RF Output | 50 ohm 1W to GND
OUT4 | 20 | Channel 4 output | Output | RF Output | 50 ohm 1W to GND
VB1 | 35 | Adjusts Channel 1 IDD current | Control | DC bias Input | connect to VB1 Output of VDD1 loopback circuit also to monitor channel S97 one per DUT | Muxed | See mux Tab
VB2 | 34 | Adjusts Channel 2 IDD current | Control | DC bias Input | connect to VB2 Output of VDD2 loopback circuit also to monitor channel S97 one per DUT | Muxed
VB3 | 9 | Adjusts Channel 3 IDD current | Control | DC bias Input | connect to VB3 Output of VDD3 loopback circuit also to monitor channel S97 one per DUT | Muxed
VB4 | 10 | Adjusts Channel 4 IDD current | Control | DC bias Input | connect to VB4 Output of VDD4 loopback circuit also to monitor channel S97 one per DUT | Muxed
VCC* | 33 | Driver power supply (for channel 1 and 2) | Power Supply | Power Supply #1 | 3.63V ±10mV (70mA) PS2 with current monitor | mon | S97-S110
VCC* | 11 | Driver power supply (for channel 3 and 4) | Power Supply | Power Supply #1 | 3.63V ±10mV (70mA) PS2 with current monitor
VDD1* | 28 | Channel 1 driver bias-T power supply | Power Supply | Power Supply #2 | VDD1 3.63V (70mA) output of loopback circuit with current monitor | Muxed | See mux Tab
VDD2* | 27 | Channel 2 driver bias-T power supply | Power Supply | Power Supply #2 | VDD1 3.63V (70mA) output of loopback circuit with current monitor | Muxed
VDD3* | 16 | Channel 3 driver bias-T power supply | Power Supply | Power Supply #2 | VDD1 3.63V (70mA) output of loopback circuit with current monitor | Muxed
VDD4* | 17 | Channel 4 driver bias-T power supply | Power Supply | Power Supply #2 | VDD1 3.63V (70mA) output of loopback circuit with current monitor | Muxed
VEA1 | 26 | Channel 1 EA modulator bias-T power supply | Power Supply | Power Supply #3 | 0 ohm to output of the VEA regulator circuit also to Banana Jack at the end of the board for -2V VEA gets regulated from PS1 40mA
VEA2 | 25 | Channel 2 EA modulator bias-T power supply | Power Supply | Power Supply #3 | 0 ohm to output of the VEA regulator circuit also to Banana Jack at the end of the board for -2V VEA gets regulated from PS1 40mA
VEA3 | 18 | Channel 3 EA modulator bias-T power supply | Power Supply | Power Supply #3 | 0 ohm to output of the VEA regulator circuit also to Banana Jack at the end of the board for -2V VEA gets regulated from PS1 40mA
VEA4 | 19 | Channel 4 EA modulator bias-T power supply | Power Supply | Power Supply #3 | 0 ohm to output of the VEA regulator circuit also to Banana Jack at the end of the board for -2V VEA gets regulated from PS1 40mA
VG1 | 32 | Channel 1 gain control | Control | DC bias Input | V5 0.2V to series 1kΩ to VG pin
VG2 | 31 | Channel 2 gain control | Control | DC bias Input | V5 0.2V to series 1kΩ to VG pin
VG3 | 12 | Channel 3 gain control | Control | DC bias Input | V5 0.2V to series 1kΩ to VG pin
VG4 | 13 | Channel 4 gain control | Control | DC bias Input | V5 0.2V to series 1kΩ to VG pin
VP1 | 30 | Channel 1 peak detector | DC Output | DC Output | Float / Data log | Muxed | See mux Tab
VP2 | 29 | Channel 2 peak detector | DC Output | DC Output | Float / Data log | Muxed
VP3 | 14 | Channel 3 peak detector | DC Output | DC Output | Float / Data log | Muxed
VP4 | 15 | Channel 4 peak detector | DC Output | DC Output | Float / Data log | Muxed
NC | 24, 36 | No connect (corner pins) | NC | NC | Float
Add test point for GND at each socket
PS1 | 3.63V, Oscillator circuit thru 2A fuse "OPAmp VDD1,2,3,4 loopback circuit thru 2A fuse, MUX thru 0 ohm, RTD, and -2V regulator thru 15A fuse
PS2 | VCC 3.63V
PS3 | VDD 3.63V
PS4 | VREF
Power up sequence
1. PS1 3.63V Oscillator circuit thru 2A fuse "OPAmp VDD1,2,3,4 loopback circuit thru 2A fuse, MUX thru 0 ohm, RTD, and -2V regulator thru 15A fuse
2. PS2 VCC
3. PS3 VDD and VG 0.2V from regulator
4. Oscillator enable channel
5. PS4 VREF
Also Connect to header at monitor end of the Board.
TP
TP
TP
TP
TP
TP
TP
TP
Also Connect to header at monitor end of the Board.

## NEG_VP_OUT Regulator (hidden)
One per Board | 15A fuse

## VP_OUT 0.2V Regulator (hidden)
One per Board | 5A fuse on Input from PS3

## Oscillator and buffer circuits (hidden)
1.Design for HTOL 14 positions per board Incal XP-160. Oscillator design is based on B1460 or B1541. One circuit can be used for two duts. as close as possible.
Route on top or bottom only with GND shield all aroud. Buffer part number changed to SY58033. Make the land pad footprints longer so I can probe them outside of the QFN device.
V4= PS3 3.63V
Do not use below
VDD1,VDD2,VDD3,VDD4 circuit after VDD current sense 4 per DUT
To Mux
VDD 1,2,3,4 goes thru active current monitor one per DUT.
To Mux
VCC goes thru active current monitor one per DUT.
See monitoring tab for details
S97-S110

## Monitoring (hidden)
Active monitor channel per DUT: | 4
|  | 1 | S97-S121
Active current monitor per DUT: | 3
VTERM
PS3 VTERM |  |  |  | IMON = Connect to S122-S147
VIN = 3.8V | 500-800mA
VDUT= VTERM 3.8V |  |  |  | QS6
RS1 | 0.05 ohm 1206 MCS1632R050FER 0.5W 1%
F1A | 0 ohm
C1A | 22uF 25V 7343 | T591D226M025ATE060
C1B | 1uF
C1C | 0.1uF
C1D | 22uF 25V 7343 | T591D226M025ATE060
C1E | 1uF
C1F | 0.1uF
PS2 |  |  |  | IMON = Connect to MUX
VIN = 2.5V-2.7V | 45mA
VDUT= VOA_X |  |  |  | QS5
RS1 | 0.5 ohm 0805 CSR0805FKR500 0.125W 1%
F1A | 0 ohm
C1A | 22uF 25V 7343 | T591D226M025ATE060
C1B | 1uF
C1C | 0.1uF
C1D | 22uF 25V 7343 | T591D226M025ATE060
C1E | 1uF
C1F | 0.1uF
PS2 |  |  |  | IMON = Connect to MUX
VIN = 2.5V-2.7V | 45mA
VDUT= VOA_Y |  |  |  | QS7
RS1 | 0.5 ohm 0805 CSR0805FKR500 0.125W 1%
F1A | 0 ohm
C1A | 22uF 25V 7343 | T591D226M025ATE060
C1B | 1uF
C1C | 0.1uF
C1D | 22uF 25V 7343 | T591D226M025ATE060
C1E | 1uF
C1F | 0.1uF
all 25 socket will also monitor temperature. Connect output header pin2 to MUX (See MUX tab)
Please add this circuit to each socket on board. A 3 pin header.Pin 1 to PS3 pin 2 Connect to monitor channel MUX
Pin 3 is Ground.
There will be one device TMP235 SOT70 per socket. Please add it to the BOM total of 25 per board. This device will not be part of the layout
for 150C use TMP235 A4 instead of TC1047A different footprint SOT70

## Edge Connector Power supplies (hidden)
PS1 | 25 sockets | 1.8V | 3.6A
RS1 | 0.01 ohm 2512 2W 1% | CSR2512A-0R01F1
F1A | 0
F1B | 0
C1A | 2200 uF 4V 7260 20% (DNI) | DNI
C1B | 470uF 4V 7343 10m ohm (DNI) | DNI
CIC | 10uF 16V 1210 X8R | C3225X8R1C106K250AB
C1D | 47uF 6.3V 1210 X7R | JMK325B7476MM-TR)
C1E | 47uF 6.3V 1210 X7R | DNI
C1F | 47uF 6.3V 1210 X7R | DNI
PS2 | 25 sockets | 2.5V | 8A
RS1 | 0.005 ohm 2512 2W 1%
F1A | 0
F1B | 0
C1A | 2200 uF 4V 7260 20% (DNI) | DNI
C1B | 470uF 4V 7343 10m ohm (DNI) | DNI
CIC | 10uF 16V 1210 X8R | C3225X8R1C106K250AB
C1D | 47uF 6.3V 1210 X7R | JMK325B7476MM-TR)
C1E | 47uF 6.3V 1210 X7R | DNI
C1F | 47uF 6.3V 1210 X7R | DNI
PS3 | 25 sockets | 3.8V | 7.5
RS1 | 0.005 ohm 2512 2W 1%
F1A | 0
F1B | 0
C1A | 2200 uF 4V 7260 20% (DNI) | DNI
C1B | 470uF 4V 7343 10m ohm (DNI) | DNI
CIC | 10uF 16V 1210 X8R | C3225X8R1C106K250AB
C1D | 47uF 6.3V 1210 X7R | JMK325B7476MM-TR)
C1E | 47uF 6.3V 1210 X7R | DNI
C1F | 47uF 6.3V 1210 X7R | DNI
PS4 | 25 sockets | 6V | 5A
RS1 | 0.02 ohm 1206 0.5W 1%
F1A | 0 |  |  |  |  | PS4 LED
F1B | 0 |  |  | PSA | PS4
C1A | 2200 uF 4V 7260 20% (DNI) | DNI |  | RL1 | 150 ohm
C1B | 470uF 4V 7343 10m ohm (DNI) | DNI |  | D1 | LED SMD
CIC | 10uF 16V 1210 X8R | C3225X8R1C106K250AB
C1D | 47uF 6.3V 1210 X7R | JMK325B7476MM-TR)
C1E | 47uF 6.3V 1210 X7R | DNI
C1F | 47uF 6.3V 1210 X7R | DNI
V3
PS2
V3

## ADC XADC Resources
ADC assignment | [5 empty] | 1 fuse for all TMUX1204 |  | V+ = VOUT8
ADC_IN00 | Current sense x4 per Socket PS3 | DUT 1 | TMUX1204DGS 1 |  | S1,S2,S3,S4
ADC_IN01 | Current sense x4 per Socket PS3 | DUT 1 | TMUX1204DGS 2 |  | S1,S2,S3,S4
ADC_IN02 | Current sense x4 per Socket PS3 | DUT 2 | TMUX1204DGS 3 |  | S1,S2,S3,S4
ADC_IN03 | Current sense x4 per Socket PS3 | DUT 2 | TMUX1204DGS 4 |  | S1,S2,S3,S4
ADC_IN04 | Current sense x4 per Socket PS3 | DUT 3 | TMUX1204DGS 5 |  | S1,S2,S3,S4
ADC_IN05 | Current sense x4 per Socket PS3 | DUT 3 | TMUX1204DGS 6 |  | S1,S2,S3,S4
ADC_IN06 | Current sense x4 per Socket PS3 | DUT 4 | TMUX1204DGS 7 |  | S1,S2,S3,S4
ADC_IN07 | Current sense x4 per Socket PS3 | DUT 4 | TMUX1204DGS 8 |  | S1,S2,S3,S4
ADC_IN08 | Current sense x4 per Socket PS3 | DUT 5 | TMUX1204DGS 9 |  | S1,S2,S3,S4
ADC_IN09 | Current sense x4 per Socket PS3 | DUT 5 | TMUX1204DGS 10 |  | S1,S2,S3,S4
ADC_IN10 | NC
ADC_IN11 | NC
ADC_IN12 | NC
ADC_IN13 | NC
ADC_IN14 | NC
ADC_IN15 | NC
XADC_0 | Case Temperature 1 Socket | DUT 1 |  | thru divider
XADC_1 | Case Temperature 1 Socket | DUT 2 |  | thru divider
XADC_2 | Case Temperature 1 Socket | DUT 3 |  | thru divider
XADC_3 | Case Temperature 1 Socket | DUT 4 |  | thru divider
XADC_4 | Case Temperature 1 Socket | DUT 5 |  | thru divider
XADC_5 | Thermistor Y2 Y14 Y18 1 socket | DUT 1 | TMUX1204DGS 11 | thru divider | S1,S2,S3
XADC_6 | Thermistor Y2 Y14 Y18 1 socket | DUT 2 | TMUX1204DGS 12 | thru divider | S1,S2,S3
XADC_7 | Thermistor Y2 Y14 Y18 1 socket | DUT 3 | TMUX1204DGS 13 | thru divider | S1,S2,S3
XADC_8 | Thermistor Y2 Y14 Y18 1 socket | DUT 4 | TMUX1204DGS 14 | thru divider | S1,S2,S3
XADC_9 | Thermistor Y2 Y14 Y18 1 socket | DUT 5 | TMUX1204DGS 15 | thru divider | S1,S2,S3
XADC_10 | NC
XADC_11 | NC
XADC_12 | NC
XADC_13 | NC
XADC_14 | NC
XADC_15 | NC
USE MIO34 MIO35 MIO36 thru 0 ohm for all 15 TMUX1204DGS MUX
| Example of X2/Y2 connection | use same for Y14 Y18 X14 X18
MIO36 for all 4 MUX1574

## Negative 1.8V
Use Module footprint From S0015

## Power Supplies
LCPS 6 0.5-5.5V | Voltage | Trim value to use | Trim resistor values in Ohms calc | Notes | Per DUT Current A | Per board Current A | Trim resistor value 23.7k x8
VOUT1 | 3.3 | 4.42k | 4444.4444 | PS1 DUT1 | 6 | 6
VOUT2 | 3.3 | 4.42k | 4444.4444 | PS1 DUT2 | 6 | 6
VOUT3 | 3.3 | 4.42k | 4444.4444 | PS1 DUT3 | 6 | 6
VOUT4 | 3.3 | 4.42k | 4444.4444 | PS1 DUT4 | 6 | 6
VOUT5 | 3.3 | 4.42k | 4444.4444 | PS1 DUT5 | 6 | 6
VOUT6 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT1 | 12 | 12
VOUT7 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT1 | 11 | 11
VOUT8 | 5 | 2.43K | 2727.2727 | Heaters, oscillator, switches supply, INA214 | 1 | 5
LCPS 7 | Voltage | Trim value to use | Trim resistor values in Ohms calc |  | Per DUT Current A | Per board Current A | Trim resistor value 54.9k x8
VOUT1 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT2 | 12 | 12
VOUT2 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT2 | 11 | 11
VOUT3 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT3 | 12 | 12
VOUT4 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT3 | 11 | 11
VOUT5 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT4 | 12 | 12
VOUT6 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT4 | 11 | 11
VOUT7 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT5 | 12 | 12
VOUT8 | 4.5 | 3.09k | 3076.9231 | Daughter card supply DUT5 | 11 | 11
HCPS 0.5-2V | Voltage
HCPS VOUT9 | 1.8V |  |  | PS3 | 0.2 | 1
HCPS VOUT10 | 2V |  |  | PS2 negative 1.8V thru regulator | 0.1 | 0.5
HCPS VOUT11 | not used
HCPS VOUT12 | not used

## Bom
Item | Qty | Ref# | Type | Value | Description | Mfr | Mfr PN | Dist PN | Board# | Cost | Total Cost

## POD
no attachment.

## Ball Map (hidden)
N/A

## Not used Reference only (hidden)
INA214 (Gain=100, positiove 26V only)
| Max Current | Rsense (ohm) | MFG | Part Number | Size (SMD) | Watts | Tol | Temp
| 25mA | 2 | Rohm | ESR10EZPF2R00 | 0805 | 0.125W | 1% | 100ppm |  |  |  |  | Caps | [7 empty] | PS3 LED
| 50mA | 1 | Rohm | MCR03EZPFL1R00 | 0805 | 0.125W | 1% | 100ppm
| 100mA | 0.5 | Stackpole | CSR0805FKR500 | 0805 | 0.125W | 1% | 200ppm |  |  |  |  | 22uF 25V 7343 | T591D226M025ATE060
| 250mA | 0.2 | Vishay | WSL0805R2000FEA | 0805 | 0.125W | 1% | 75ppm | [9 empty] | PSA | V++
| 500mA | 0.1 | Stackpole | CSR1206FKR100 | 1206 | 0.5W | 1% | 100ppm | [9 empty] | PSB | PS3
| 1A | 0.05 | Ohmite | MCS1632R050FER | 1206 | 0.5W | 1% | 50ppm | [9 empty] | RL2 | 4.7K
| 2.5A | 0.02 | Ohmite | MCS1632R020FER | 1206 | 0.5W | 1% | 50ppm |  | For 15 amps fuse, let’s use ED5019-ND. The holes on the PCB will need to be bigger as well. Like the 20amp | [7 empty] | RL1 | 22.1 ohm
| 5A | 0.01 | Riedon | CSR2512A-0R01F1 | 2512 | 2W | 1% | 50ppm | [9 empty] | D1 | LED SMD
| 10A | 0.005 | TT Elctronics | LRF3W-R005JW | 2512 | 2W | 1% | 50ppm | [9 empty] | Q1 | MMBT3904
| 25A | 0.002 | Stackpole | CSNL2512 FT2L00 | 2512 | 2W | 1% | 50ppm
| 50A | 0.001 | Riedon | CSR2512C-0R001F1 | 2512 | 3W | 1% | 50ppm
| 100A | 0.0005 | Stackpole | CSS2725FTL500 | 2725 | 4W | 1% | 50ppm
PS4 LED
[19 empty] | PSA | PS4
[19 empty] | RL1 | 11 ohm
[19 empty] | D1 | LED SMD
available R
------ | ++++ | [7 empty] | 16
R1 | R2 | [7 empty] | 22
392 | 866 |  | Vout = | 4.0115 | <-- NEED | 0.75 |  |  | 15
28
36
R1 | R2 | [7 empty] | 37 |  |  | 0.75 | 1.05 | 0.7875
392 | 392 |  | Vout = | 2.5 | <-- NEED | 2.5 | [5 empty] | 1.2 | 1.05 | 1.26
[9 empty] | 300 |  |  | 1.8 | 1.05 | 1.89
[9 empty] | 250 |  |  | 2.5 | 1.05 | 2.625
R1 | R2 | [7 empty] | 220 |  |  | 3.3 | 1.05 | 3.465
200 | 681 |  | Vout = | 5.5062 | <-- NEED |  |  |  | 390
487
511
R1 | R2 | [7 empty] | 576
200 | 680 |  | Vout = | 5.5 | <-- NEED | 3.3 |  |  | 1000
R1 | R2
198 | 673 |  | Vout = | 5.4987 | <-- NEED | 3.3
[10 empty] | ps4 |  | ps5 |  | ps6
[10 empty] | R1 | R2 | R1 | R2 | R1 | R2
[16 empty] | 1V8 regulated from PS3 | LT1085CM
VIN = PS3
VOUT = 1V8
R1 = 392 ohm 1%
R2 = removable resistor 174 ohm 1%
[16 empty] | 2V5 regulated from PS3 | LT1085CM
VIN = PS3
VOUT = 2V5
R1 = 392 ohm 1%
R2 = removable resistor 392 ohm 1%